## Table of Contents
Reference measurements for model 130: Premium System phantom with serial number 130-0062
This Excel Spreadsheet for your 130-0062 phantom is provided by CaliberMRI and may reference other documents.
bottom of each data sheet. Please contact CaliberMRI if you have any questions.
https://qmri.com/contact/
Worksheet | Contents
NiCl Solutions @ 1.5T | T1 Solution (NiCl2) Reference Measurements at 1.5T
NiCl Solutions @ 3.0T | T1 Solution (NiCl2) Reference Measurements at 3.0T
MnCl Solutions @ 1.5T | T2 Solution (MnCl2) Reference Measurements at 1.5T
MnCl Solutions @ 3.0T | T2 Solution (MnCl2) Reference Measurements at 3.0T
PD Solutions @ 3.0T | PD Reference Measurements at 3.0T
CuSO4 Solutions @ 3.0T | CuSO4 Reference Measurements at 3.0T
CMRI LC Values | CMRI LC Thermometer Transition Temperatures
Uncertainty values represent one standard deviation of variation, though the original measurement report may use
three standard deviations instead. Methods of determining uncertainty vary depending on the measurement source. See
the relevant measurement report for more details.
Preliminary stability data from NIST and CaliberMRI indicates that mimic solutions remain stable for three years.
CaliberMRI recommends replacing your phantom after three years.

## NiCl Solutions @ 1.5T
T1 Solution (NiCl2) Reference Measurements at 1.5T for serial number 130-0062
Solution | Temperature (C) | Concentration (mM) | T1 (ms) | T1 Uncertainty (ms) | T2 (ms) | T2 Uncertainty (ms) | Notes
T1-1 | 20 | 0.29 | 1741 |  | 1540 |  | a
T1-2 | 20 | 0.6 | 1270 |  | 1130 |  | a
T1-3 | 20 | 1.04 | 962 |  | 845 |  | a
T1-4 | 20 | 1.64 | 685 |  | 617 |  | a
T1-5 | 20 | 2.52 | 487 |  | 430 |  | a
T1-6 | 20 | 3.68 | 346 |  | 306 |  | a
T1-7 | 20 | 5.43 | 246 |  | 218 |  | a
T1-8 | 20 | 7.74 | 179 |  | 159 |  | a
T1-9 | 20 | 11.3 | 124 |  | 110 |  | a
T1-10 | 20 | 16.5 | 88 |  | 77 |  | a
T1-11 | 20 | 23.3 | 62 |  | 54 |  | a
T1-12 | 20 | 32.7 | 44 |  | 39 |  | a
T1-13 | 20 | 46 | 31 |  | 28 |  | a
T1-14 | 20 | 65.3 | 22 |  | 20 |  | a
[a] The derivation of T1 and T2 values at 20°C by CaliberMRI using NIST methodology is based on 1) NIST T1 and T2 measurements at 1.5T and 3T on a different batch NiCl2 and MnCl2 solutions with well-known concentrations, measured by ICP-MS/OES technique, and 2) NIST traceable T1 and T2 measurements at 3T of your phantom's solutions, as per the report at (link below)
Report: | https://qcalsoftware.com/measurements/nist_s2

## NiCl Solutions @ 3.0T
T1 Solution (NiCl2) Reference Measurements at 3.0T for serial number 130-0062
Solution | Temperature (C) | Concentration (mM) | T1 (ms) | T1 Uncertainty (ms) | T2 (ms) | T2 Uncertainty (ms) | Notes
T1-1 | 16 | 0.29 | 1766.7 | 9.51 | 1348.9 | 6.32 | a, b
T1-1 | 18 | 0.29 | 1830.3 | 9.81 | 1455.1 | 6.57 | a, b
T1-1 | 20 | 0.29 | 1884 | 10.11 | 1489.4 | 6.8 | a, b
T1-1 | 22 | 0.29 | 1937.3 | 9.51 | 1520.8 | 6.32 | a, b
T1-1 | 24 | 0.29 | 1987.5 | 9.81 | 1532.3 | 6.57 | a, b
T1-1 | 26 | 0.29 | 2066.9 | 10.11 | 1578.6 | 6.8 | a, b
T1-2 | 16 | 0.6 | 1274.1 | 6.48 | 990.8 | 4.29 | a, b
T1-2 | 18 | 0.6 | 1317.7 | 6.66 | 1010.3 | 4.33 | a, b
T1-2 | 20 | 0.6 | 1330.2 | 6.8 | 1026.8 | 4.34 | a, b
T1-2 | 22 | 0.6 | 1355.3 | 6.48 | 1040.2 | 4.29 | a, b
T1-2 | 24 | 0.6 | 1367.8 | 6.66 | 1048.1 | 4.33 | a, b
T1-2 | 26 | 0.6 | 1395.9 | 6.8 | 1067.1 | 4.34 | a, b
T1-3 | 16 | 1.04 | 950.7 | 4.61 | 731.7 | 3.03 | a, b
T1-3 | 18 | 1.04 | 963.6 | 4.67 | 735.3 | 3.06 | a, b
T1-3 | 20 | 1.04 | 987.3 | 4.74 | 737 | 3.11 | a, b
T1-3 | 22 | 1.04 | 1000.8 | 4.61 | 744.4 | 3.03 | a, b
T1-3 | 24 | 1.04 | 1015.7 | 4.67 | 755.9 | 3.06 | a, b
T1-3 | 26 | 1.04 | 1030.8 | 4.74 | 768.7 | 3.11 | a, b
T1-4 | 16 | 1.64 | 675.1 | 3.31 | 519.3 | 2.15 | a, b
T1-4 | 18 | 1.64 | 686.9 | 3.34 | 520.9 | 2.16 | a, b
T1-4 | 20 | 1.64 | 690.1 | 3.37 | 521.3 | 2.19 | a, b
T1-4 | 22 | 1.64 | 695 | 3.31 | 523.1 | 2.15 | a, b
T1-4 | 24 | 1.64 | 701.1 | 3.34 | 528.5 | 2.16 | a, b
T1-4 | 26 | 1.64 | 709.5 | 3.37 | 538.2 | 2.19 | a, b
T1-5 | 16 | 2.52 | 483.9 | 2.32 | 358.9 | 1.48 | a, b
T1-5 | 18 | 2.52 | 482.9 | 2.34 | 355.1 | 1.49 | a, b
T1-5 | 20 | 2.52 | 485 | 2.35 | 354.4 | 1.5 | a, b
T1-5 | 22 | 2.52 | 486.9 | 2.32 | 356.1 | 1.48 | a, b
T1-5 | 24 | 2.52 | 490.2 | 2.34 | 359.8 | 1.49 | a, b
T1-5 | 26 | 2.52 | 494.6 | 2.35 | 365.2 | 1.5 | a, b
T1-6 | 16 | 3.68 | 343 | 1.64 | 252.2 | 1.03 | a, b
T1-6 | 18 | 3.68 | 341.5 | 1.65 | 249.2 | 1.04 | a, b
T1-6 | 20 | 3.68 | 341.6 | 1.66 | 248 | 1.05 | a, b
T1-6 | 22 | 3.68 | 342.6 | 1.64 | 248.2 | 1.03 | a, b
T1-6 | 24 | 3.68 | 344.2 | 1.65 | 250.1 | 1.04 | a, b
T1-6 | 26 | 3.68 | 346.7 | 1.66 | 253.8 | 1.05 | a, b
T1-7 | 16 | 5.43 | 243.8 | 1.16 | 178.7 | 0.73 | a, b
T1-7 | 18 | 5.43 | 241.8 | 1.16 | 176.1 | 0.73 | a, b
T1-7 | 20 | 5.43 | 240.9 | 1.17 | 175 | 0.73 | a, b
T1-7 | 22 | 5.43 | 240.8 | 1.16 | 175.1 | 0.73 | a, b
T1-7 | 24 | 5.43 | 241.3 | 1.16 | 176.2 | 0.73 | a, b
T1-7 | 26 | 5.43 | 242.4 | 1.17 | 178.1 | 0.73 | a, b
T1-8 | 16 | 7.74 | 177.7 | 0.82 | 129.8 | 0.51 | a, b
T1-8 | 18 | 7.74 | 175.9 | 0.82 | 127.3 | 0.51 | a, b
T1-8 | 20 | 7.74 | 174.9 | 0.83 | 126.3 | 0.52 | a, b
T1-8 | 22 | 7.74 | 174.6 | 0.82 | 126.5 | 0.51 | a, b
T1-8 | 24 | 7.74 | 174.8 | 0.82 | 126.9 | 0.51 | a, b
T1-8 | 26 | 7.74 | 175.5 | 0.83 | 128.3 | 0.52 | a, b
T1-9 | 16 | 11.3 | 123 | 0.58 | 89 | 0.36 | a, b
T1-9 | 18 | 11.3 | 121.8 | 0.58 | 87.7 | 0.36 | a, b
T1-9 | 20 | 11.3 | 121.1 | 0.58 | 86.9 | 0.36 | a, b
T1-9 | 22 | 11.3 | 120.8 | 0.58 | 86.8 | 0.36 | a, b
T1-9 | 24 | 11.3 | 120.9 | 0.58 | 87.3 | 0.36 | a, b
T1-9 | 26 | 11.3 | 121.3 | 0.58 | 88.4 | 0.36 | a, b
T1-10 | 16 | 16.5 | 87.5 | 0.41 | 62.4 | 0.25 | a, b
T1-10 | 18 | 16.5 | 86.4 | 0.41 | 61.6 | 0.25 | a, b
T1-10 | 20 | 16.5 | 85.8 | 0.41 | 60.9 | 0.26 | a, b
T1-10 | 22 | 16.5 | 85 | 0.41 | 61 | 0.25 | a, b
T1-10 | 24 | 16.5 | 85 | 0.41 | 61.3 | 0.25 | a, b
T1-10 | 26 | 16.5 | 85.3 | 0.41 | 62.1 | 0.26 | a, b
T1-11 | 16 | 23.3 | 61.5 | 0.29 | 43.6 | 0.18 | a, b
T1-11 | 18 | 23.3 | 60.7 | 0.29 | 42.9 | 0.18 | a, b
T1-11 | 20 | 23.3 | 60.2 | 0.29 | 42.4 | 0.18 | a, b
T1-11 | 22 | 23.3 | 60 | 0.29 | 42.4 | 0.18 | a, b
T1-11 | 24 | 23.3 | 60 | 0.29 | 42.7 | 0.18 | a, b
T1-11 | 26 | 23.3 | 60.2 | 0.29 | 43.2 | 0.18 | a, b
T1-12 | 16 | 32.7 | 43.8 | 0.15 | 31.6 | 0.09 | a, b
T1-12 | 18 | 32.7 | 43.2 | 0.15 | 31 | 0.09 | a, b
T1-12 | 20 | 32.7 | 42.9 | 0.15 | 30.7 | 0.09 | a, b
T1-12 | 22 | 32.7 | 42.7 | 0.15 | 30.6 | 0.09 | a, b
T1-12 | 24 | 32.7 | 42.7 | 0.15 | 30.8 | 0.09 | a, b
T1-12 | 26 | 32.7 | 42.8 | 0.15 | 30.7 | 0.09 | a, b
T1-13 | 16 | 46 | 31.1 | 0.21 | 22.4 | 0.13 | a, b
T1-13 | 18 | 46 | 30.6 | 0.21 | 22 | 0.13 | a, b
T1-13 | 20 | 46 | 30.4 | 0.21 | 21.8 | 0.13 | a, b
T1-13 | 22 | 46 | 30.3 | 0.21 | 21.7 | 0.13 | a, b
T1-13 | 24 | 46 | 30.2 | 0.21 | 21.8 | 0.13 | a, b
T1-13 | 26 | 46 | 30.3 | 0.21 | 22.1 | 0.13 | a, b
T1-14 | 16 | 65.3 | 21.9 | 0.1 | 15.8 | 0.06 | a, b
T1-14 | 18 | 65.3 | 21.6 | 0.1 | 15.5 | 0.06 | a, b
T1-14 | 20 | 65.3 | 21.4 | 0.1 | 15.3 | 0.06 | a, b
T1-14 | 22 | 65.3 | 21.3 | 0.1 | 15.2 | 0.06 | a, b
T1-14 | 24 | 65.3 | 21.3 | 0.1 | 15.3 | 0.06 | a, b
T1-14 | 26 | 65.3 | 21.3 | 0.1 | 15.5 | 0.06 | a, b
[a] NIST Traceable measurement performed at NIST Boulder. Report: (link below)
Report: | https://qcalsoftware.com/measurements/nist_s2
[b] Uncertainty values represent one standard deviation of variation, though the original measurement report may use three standard deviations instead. Methods of determining uncertainty vary depending on the measurement source. See the relevant measurement report for more details.

## MnCl Solutions @ 1.5T
T2 Solution (MnCl2) Reference Measurements at 1.5T for serial number 130-0062
Solution | Temperature (C) | Concentration (mM) | T1 (ms) | T1 Uncertainty (ms) | T2 (ms) | T2 Uncertainty (ms) | Notes
T2-2 | 20 | 0.0181 | 1898 |  | 493 |  | a, b
T2-3 | 20 | 0.0282 | 1655 |  | 355 |  | a, b
T2-4 | 20 | 0.0434 | 1353 |  | 237 |  | a, b
T2-6 | 20 | 0.0934 | 900 |  | 123 |  | a, b
T2-7 | 20 | 0.1353 | 707 |  | 88 |  | a, b
T2-8 | 20 | 0.193 | 528 |  | 62 |  | a, b
T2-9 | 20 | 0.2768 | 380 |  | 42 |  | a, b
T2-10 | 20 | 0.4276 | 259 |  | 27 |  | a, b
T2-11 | 20 | 0.5555 | 200 |  | 21 |  | a, b
T2-12 | 20 | 0.7902 | 140 |  | 15 |  | a, b
T2-13 | 20 | 1.1274 | 103 |  | 10 |  | a, b
T2-14 | 20 | 1.5996 | 73 |  | 7 |  | a, b
[a] The derivation of T1 and T2 values at 20°C by CaliberMRI using NIST methodology is based on 1) NIST T1 and T2 measurements at 1.5T and 3T on a different batch NiCl2 and MnCl2 solutions with well-known concentrations, measured by ICP-MS/OES technique, and 2) NIST traceable T1 and T2 measurements at 3T of your phantom's solutions, as per the report at (link below)
Report: | https://qcalsoftware.com/measurements/nist_s2
[b] T2-1 and T2-5 of your (SN130-0043-130-0132) solutions should not be used for reference data or research purposes. Please contact CaliberMRI directly with any questions at support@qmri.com

## MnCl Solutions @ 3.0T
T2 Solution (MnCl2) Reference Measurements at 3.0T for serial number 130-0062
Solution | Temperature (C) | Concentration (mM) | T1 (ms) | T1 Uncertainty (ms) | T2 (ms) | T2 Uncertainty (ms) | Notes
T2-2 | 16 | 0.018 | 1875 | 15.62 | 357.4 | 1.12 | a, b, c
T2-2 | 18 | 0.018 | 2073.9 | 17.28 | 370.5 | 1.16 | a, b, c
T2-2 | 20 | 0.018 | 2185.5 | 18.21 | 379.5 | 1.19 | a, b, c
T2-2 | 22 | 0.018 | 2302.8 | 19.19 | 392.5 | 1.23 | a, b, c
T2-2 | 24 | 0.018 | 2424.6 | 20.21 | 406 | 1.27 | a, b, c
T2-2 | 26 | 0.018 | 2546 | 21.22 | 420.8 | 1.32 | a, b, c
T2-3 | 16 | 0.028 | 1657.7 | 13.81 | 247.9 | 0.78 | a, b, c
T2-3 | 18 | 0.028 | 1783.2 | 14.86 | 257.5 | 0.81 | a, b, c
T2-3 | 20 | 0.028 | 1901.3 | 15.84 | 267.3 | 0.84 | a, b, c
T2-3 | 22 | 0.028 | 2009 | 16.74 | 277 | 0.87 | a, b, c
T2-3 | 24 | 0.028 | 2119.3 | 17.66 | 287.5 | 0.9 | a, b, c
T2-3 | 26 | 0.028 | 2232.4 | 18.6 | 299.1 | 0.94 | a, b, c
T2-4 | 16 | 0.043 | 1341.1 | 11.18 | 161.1 | 0.5 | a, b, c
T2-4 | 18 | 0.043 | 1466.8 | 12.22 | 169 | 0.53 | a, b, c
T2-4 | 20 | 0.043 | 1550 | 12.92 | 175.1 | 0.55 | a, b, c
T2-4 | 22 | 0.043 | 1626.5 | 13.55 | 181.5 | 0.57 | a, b, c
T2-4 | 24 | 0.043 | 1715.1 | 14.29 | 187.2 | 0.59 | a, b, c
T2-4 | 26 | 0.043 | 1801.6 | 15.01 | 192.2 | 0.6 | a, b, c
T2-6 | 16 | 0.093 | 892.3 | 7.44 | 81.9 | 0.26 | a, b, c
T2-6 | 18 | 0.093 | 972.5 | 8.1 | 85.9 | 0.27 | a, b, c
T2-6 | 20 | 0.093 | 1026.4 | 8.55 | 88.9 | 0.28 | a, b, c
T2-6 | 22 | 0.093 | 1083.4 | 9.03 | 92.3 | 0.29 | a, b, c
T2-6 | 24 | 0.093 | 1141.5 | 9.51 | 96 | 0.3 | a, b, c
T2-6 | 26 | 0.093 | 1202.4 | 10.02 | 99.9 | 0.31 | a, b, c
T2-7 | 16 | 0.135 | 711.1 | 5.93 | 59.5 | 0.19 | a, b, c
T2-7 | 18 | 0.135 | 757.3 | 6.31 | 61.3 | 0.19 | a, b, c
T2-7 | 20 | 0.135 | 805.1 | 6.71 | 63.4 | 0.2 | a, b, c
T2-7 | 22 | 0.135 | 848.6 | 7.11 | 65.7 | 0.21 | a, b, c
T2-7 | 24 | 0.135 | 892.4 | 7.44 | 67.8 | 0.21 | a, b, c
T2-7 | 26 | 0.135 | 937.1 | 7.81 | 70.6 | 0.22 | a, b, c
T2-8 | 16 | 0.193 | 531.6 | 4.43 | 41.5 | 0.13 | a, b, c
T2-8 | 18 | 0.193 | 563.9 | 4.7 | 42.6 | 0.13 | a, b, c
T2-8 | 20 | 0.193 | 600 | 5 | 44.2 | 0.14 | a, b, c
T2-8 | 22 | 0.193 | 634.6 | 5.29 | 45.9 | 0.14 | a, b, c
T2-8 | 24 | 0.193 | 669.6 | 5.58 | 47.7 | 0.15 | a, b, c
T2-8 | 26 | 0.193 | 703.5 | 5.86 | 49.3 | 0.15 | a, b, c
T2-9 | 16 | 0.277 | 385 | 3.21 | 28 | 0.09 | a, b, c
T2-9 | 18 | 0.277 | 408.1 | 3.4 | 28.9 | 0.09 | a, b, c
T2-9 | 20 | 0.277 | 431.2 | 3.59 | 29.9 | 0.09 | a, b, c
T2-9 | 22 | 0.277 | 454.8 | 3.79 | 30.9 | 0.1 | a, b, c
T2-9 | 24 | 0.277 | 478.8 | 3.99 | 32.1 | 0.1 | a, b, c
T2-9 | 26 | 0.277 | 503.3 | 4.19 | 33.2 | 0.1 | a, b, c
T2-10 | 16 | 0.428 | 262.1 | 2.18 | 18.2 | 0.06 | a, b, c
T2-10 | 18 | 0.428 | 277.4 | 2.31 | 18.8 | 0.06 | a, b, c
T2-10 | 20 | 0.428 | 292.9 | 2.44 | 19.4 | 0.06 | a, b, c
T2-10 | 22 | 0.428 | 308.9 | 2.57 | 20.1 | 0.06 | a, b, c
T2-10 | 24 | 0.428 | 325.2 | 2.71 | 20.8 | 0.07 | a, b, c
T2-10 | 26 | 0.428 | 341.7 | 2.85 | 21.6 | 0.07 | a, b, c
T2-11 | 16 | 0.555 | 203.5 | 1.7 | 13.9 | 0.04 | a, b, c
T2-11 | 18 | 0.555 | 214.8 | 1.79 | 14.3 | 0.04 | a, b, c
T2-11 | 20 | 0.555 | 226.5 | 1.89 | 14.7 | 0.05 | a, b, c
T2-11 | 22 | 0.555 | 238.3 | 1.99 | 15.2 | 0.05 | a, b, c
T2-11 | 24 | 0.555 | 250.6 | 2.09 | 15.8 | 0.05 | a, b, c
T2-11 | 26 | 0.555 | 263.1 | 2.19 | 16.4 | 0.05 | a, b, c
T2-12 | 16 | 0.79 | 141.2 | 1.18 | 9.9 | 0.03 | a, b, c
T2-12 | 18 | 0.79 | 149.3 | 1.24 | 10.2 | 0.03 | a, b, c
T2-12 | 20 | 0.79 | 158.2 | 1.32 | 10.5 | 0.03 | a, b, c
T2-12 | 22 | 0.79 | 166.6 | 1.39 | 10.9 | 0.03 | a, b, c
T2-12 | 24 | 0.79 | 175.3 | 1.46 | 11.2 | 0.04 | a, b, c
T2-12 | 26 | 0.79 | 184.1 | 1.53 | 11.7 | 0.04 | a, b, c
T2-13 | 16 | 1.127 | 104.6 | 0.87 | 6.8 | 0.02 | a, b, c
T2-13 | 18 | 1.127 | 110.5 | 0.92 | 7 | 0.02 | a, b, c
T2-13 | 20 | 1.127 | 116.7 | 0.97 | 7.3 | 0.02 | a, b, c
T2-13 | 22 | 1.127 | 123.2 | 1.03 | 7.5 | 0.02 | a, b, c
T2-13 | 24 | 1.127 | 129.7 | 1.08 | 7.8 | 0.02 | a, b, c
T2-13 | 26 | 1.127 | 136.3 | 1.14 | 8.1 | 0.03 | a, b, c
T2-14 | 16 | 1.6 | 73.7 | 0.61 | 4.8 | 0.02 | a, b, c
T2-14 | 18 | 1.6 | 77.9 | 0.65 | 4.9 | 0.02 | a, b, c
T2-14 | 20 | 1.6 | 82.5 | 0.69 | 5.1 | 0.02 | a, b, c
T2-14 | 22 | 1.6 | 87.1 | 0.73 | 5.3 | 0.02 | a, b, c
T2-14 | 24 | 1.6 | 91.8 | 0.77 | 5.5 | 0.02 | a, b, c
T2-14 | 26 | 1.6 | 96.5 | 0.8 | 5.7 | 0.02 | a, b, c
[a] NIST Traceable measurement performed at NIST Boulder. Report: (link below)
Report: | https://qcalsoftware.com/measurements/nist_s2
[b] Uncertainty values represent one standard deviation of variation, though the original measurement report may use three standard deviations instead. Methods of determining uncertainty vary depending on the measurement source. See the relevant measurement report for more details.
[c] T2-1 and T2-5 of your (SN130-0043-130-0132) solutions should not be used for reference data or research purposes. Please contact CaliberMRI directly with any questions at support@qmri.com

## PD Solutions @ 3.0T
PD Reference Measurements at 3.0T for serial number 130-0062
Solution | Temperature (C) | Concentration (mM) | PD (%) | PD Uncertainty (%) | Notes
H2O_5 | 20 |  | 5.1 | 0.06 | a, b
H2O_10 | 20 |  | 10.2 | 0.12 | a, b
H2O_15 | 20 |  | 15 | 0.18 | a, b
H2O_20 | 20 |  | 20.7 | 0.25 | a, b
H2O_25 | 20 |  | 25.5 | 0.31 | a, b
H2O_30 | 20 |  | 30 | 0.36 | a, b
H2O_35 | 20 |  | 34.9 | 0.42 | a, b
H2O_40 | 20 |  | 39.8 | 0.48 | a, b
H2O_50 | 20 |  | 48.4 | 0.58 | a, b
H2O_60 | 20 |  | 59.1 | 0.71 | a, b
H2O_70 | 20 |  | 68.5 | 0.83 | a, b
H2O_80 | 20 |  | 78.3 | 0.94 | a, b
H2O_90 | 20 |  | 90.6 | 1.09 | a, b
H2O_100 | 20 |  | 99.5 | 1.2 | a, b
[a] Measurement performed at NIST Boulder (not certified) on a representative sample but not from the batch used in this phantom. NIST Report: (link below)
Report: | https://qcalsoftware.com/measurements/nist_s2.5
[b] Uncertainty values represent one standard deviation of variation, though the original measurement report may use three standard deviations instead. Methods of determining uncertainty vary depending on the measurement source. See the relevant measurement report for more details.

## CuSO4 Solutions @ 3.0T
CuSO4 Reference Measurements at 3.0T for serial number 130-0062
Solution | Temperature (C) | Concentration (mM) | T1 (ms) | T1 Uncertainty (ms) | T2 (ms) | T2 Uncertainty (ms) | Notes
CuSO4 | 16 | 2.9 | 350.2 | 2.57 | 300.3 | 2.2 | a, b, c
CuSO4 | 18 | 2.9 | 371 | 2.72 | 317.9 | 2.33 | a, b, c
CuSO4 | 20 | 2.9 | 391.5 | 2.87 | 335.6 | 2.46 | a, b, c
CuSO4 | 22 | 2.9 | 412 | 3.02 | 352.5 | 2.59 | a, b, c
CuSO4 | 24 | 2.9 | 432.8 | 3.17 | 370.6 | 2.72 | a, b, c
CuSO4 | 26 | 2.9 | 453.8 | 3.33 | 388.9 | 2.85 | a, b, c
[a] Please note that CuSO4 solution should not be used as a tissue mimic. The listed CuSO4 concentration and relaxation properties are for guidance only. The actual concentration may differ slightly.
[b] Measurement performed at NIST Boulder (not certified) on a representative sample but not from the batch used in this phantom.
[c] Uncertainty values represent one standard deviation of variation, though the original measurement report may use three standard deviations instead. Methods of determining uncertainty vary depending on the measurement source.

## CMRI LC Values
CMRI LC Thermometer Transition Temperatures for serial number 130-0062
Solution | Measured Transition Temperature (C) | Transition Temperature Uncertainty (C)
LC-1 | 15.5 | 0.3
LC-2 | 16.5 | 0.3
LC-3 | 17.5 | 0.3
LC-4 | 18.8 | 0.3
LC-5 | 19.8 | 0.4
LC-6 | 20.9 | 0.5
LC-7 | 21.9 | 0.5
LC-8 | 23.4 | 0.5
LC-9 | 24.1 | 0.6
LC-10 | 25 | 0.6